## Slide 1
Testing how far exposure can be reduced
A known-truth water / plexiglass simulation for Brillouin spectroscopy
Noise physics → BFS fitting → clustering → regional averaging
100 s
to
0.004 s
Question
Can spatial redundancy recover spectra after individual-pixel fitting fails?
Brillouin spectroscopy noise-tolerance study
Notes: Frame the study as a method-development question, not a claim about the final biological limit.
[Sources]
- Local experiment output: outputs/simulated_water_plexiglass/generated_four_stage_noise_grid.png

## Slide 2
EXPERIMENTAL DESIGN
A known boundary makes every recovery step measurable
Cell Builder truth
100% WATER
100% PLEXIGLASS
30 × 30 pixels • 1 μm step • exact truth sidecar
Initial exposure sweep
100 s
1.00× noise
30 s
1.83× noise
10 s
3.16× noise
3 s
5.77× noise
1 s
10.0× noise
0.3 s
18.3× noise
0.1 s
31.6× noise
t
=
100
m²
Signal ∝ exposure
Relative shot noise ∝ 1/√t
02
Notes: The boundary is intentionally simple: the goal is to isolate noise tolerance before returning to a cell geometry.
[Sources]
- simulator/generate_water_plexiglass_sweep.py
- simulator/resources/materials.json
- outputs/simulated_water_plexiglass/*_truth.npz

## Slide 3
HOW NOISE IS CALCULATED
The simulator separates physical line broadening from detector noise
01
Thermal acoustic field
02
Laser phase noise
03
Detector chain
Acoustic density follows an Ornstein–Uhlenbeck process.
Thermal fluctuations average over many coherence times; at realistic exposure they become negligible speckle.
φ(t) = √ΔΩ · W(t)
A Brownian phase walk produces Lorentzian laser broadening. Widths add analytically to the material linewidth.
Poisson shot noise
+ dark counts
+ read noise
+ EMCCD excess noise (optional)
Exposure mainly changes detected Brillouin photons—and therefore the visible pixel-to-pixel scatter.
03
Notes: Distinguish linewidth physics from visible detector scatter. Thermal noise sets the line; shot noise dominates these camera-like spectra.
[Sources]
- simulator/core/noise.py
- simulator/core/spectrum.py
- Oscar Nieves, Noise in Stimulated Brillouin Scattering, PhD thesis (2021), equations 3.12, 3.18, 3.23, 3.27, 5.13 and 5.14

## Slide 4
BFS FITTING
At 0.1 s BFS is noisier, but no longer collapses toward zero
Lorentzian model
I(ν) = A / [1 + ((ν − νB)/γ)²] + C
Fit parameters
A amplitude
νB Brillouin frequency shift
γ half-width at half maximum
C baseline
Corrected short-exposure behaviour
•
Only 0.1% of the 100 s photon budget remains.
•
Elastic and Brillouin photons now both scale linearly with exposure.
•
Read noise remains fixed per acquisition, so pixel-level BFS estimates become visibly less precise.
•
The two fitted regions remain centred near 5.51 and 11.75 GHz rather than collapsing to the lower bound.
The earlier low-BFS expansion was a fixed-elastic-peak simulation artefact.
04
Notes: Emphasise that the plotting range exposes fit failure. It is not a rendering bug and not proof that all spatial information has vanished.
[Sources]
- brillouin_fitting.py: lorentzian and fit_lorentzian_iterative
- simulator/core/scan.py: fit_scan
- outputs/simulated_water_plexiglass/half_water_half_plexiglass_noise_31.6228x.csv

## Slide 5
GROUPING MATERIAL REGIONS
Clustering uses spectral shape, so it survives after single-pixel BFS fails
01
Fold
Average Stokes and anti-Stokes sides
02
Trim
Remove Rayleigh tail; retain ≤13.3 GHz
03
Normalise
Subtract baseline and L1-normalise each pixel
04
Cluster
K-means with k = 2 on the full spectrum
05
Match truth
Resolve label swapping against known material truth
Why it works
Water and plexiglass peaks are separated by ~6.24 GHz. Even when each peak centre is too noisy to fit precisely, the complete spectral shape still carries enough information to separate the two groups.
05
Notes: This is deliberately an easy two-material case. The next experiment should reduce spectral separation and add mixed regions.
[Sources]
- analysis/generated_water_plexiglass_grid.py
- brillouin_preprocessing.py
- scikit-learn KMeans algorithm used locally

## Slide 6
REGENERATING THE CURVES
Regional averaging reconstructs the peaks from many weak measurements
Average all spectra assigned to a region → independent noise falls approximately as 1/√N → refit or inspect the recovered material curve
06
Notes: Read the figure by rows. At 0.1 s the pixel BFS map collapses, but clustering still finds the boundary and the averaged peaks remain visible.
[Sources]
- outputs/simulated_water_plexiglass/generated_four_stage_noise_grid.png
- analysis/generated_water_plexiglass_grid.py

## Slide 7
WHY THIS IS PRACTICAL
Regional recovery could trade exposure for throughput
If the information can be accurately reproduced after grouping and averaging:
01
Reduce exposure time per pixel without immediately losing material-level spectra.
02
Generate more samples—or scan larger fields—in the same instrument time.
03
Use known-truth simulations to choose an exposure before committing biological samples.
04
Define acceptance thresholds for region accuracy, BFS error and Lorentzian fit quality.
05
Validate the selected exposure on real repeated scans before adopting it experimentally.
4–5 ms
observed boundary in this binary test
Next decision
What is the shortest exposure that remains reliable for less-separated biological materials?
07
Notes: End with a proposal: retain the workflow, but move from a binary easy case to progressively harder material separations and mixed boundaries.
[Sources]
- outputs/simulated_water_plexiglass/manifest.csv
- outputs/simulated_water_plexiglass/generated_four_stage_noise_grid.png

## Slide 8
PARAMETER AUDIT • 1 OF 2
Every stochastic and acquisition control used in the limit test
Parameter
Value used
Where it enters
Status
Exposure time
0.030 → 0.004 s
Signal + dark accumulation
Test variable
Temperature
300 K (26.85 °C)
Thermal signal scaling
Plausible ambient
Thermal speckle modes
Auto: t / coherence time
Gamma-distributed OU power
Physics-derived
BFS jitter / slow drift
0 MHz rms
Per-pixel common shift
Disabled to isolate SNR
Random seeds
20260824 + level index
All stochastic draws
Recorded / reproducible
PSF FWHM
0.60 μm
Boundary mixing
Plausible confocal spot
Grid / step
30 × 30 / 1 μm
900 spectra per scan
Experimental design
Temperature scaling
Enabled
Material peak area ∝ T
Thermal model
Conserve peak area
False
Material efficiencies retained
Realistic brightness contrast
Drift = 0 is an intentional controlled-test assumption—not a claim that the real laser has zero drift.
08
Notes: The limit test isolates photon statistics. A later robustness study should sweep measured laser drift rather than inventing it.
[Sources]
- simulator/core/scan.py
- simulator/core/noise.py
- simulator/generate_water_plexiglass_sweep.py
- NIST temperature reference: https://www.nist.gov/pml/owm/si-units-temperature

## Slide 9
PARAMETER AUDIT • 2 OF 2
Optical response and detector values are explicit—and not equally validated
Optical and spectral response
Parameter
Value
Evidence status
Laser wavelength
720 nm
Back-solved assumption
Scattering angle
180°
Backscatter geometry
Laser linewidth
0.10 MHz FWHM
Plausible narrow laser
Spectrometer Lorentzian
0.040 GHz FWHM
Model choice
Spectrometer Gaussian
0.400 GHz FWHM
Fit to local line shape
Elastic intrinsic width
0.030 GHz FWHM
Model choice
Frequency axis
−14.9 to +14.9 GHz
Matched local CSVs
Fit markers
3.0 to 13.2 GHz
Expanded for PMMA
Photon and detector chain
Parameter
Value
Evidence status
Brillouin photon rate
1900 counts s⁻¹
Calibrated to local scans
Elastic peak
30,000 counts @ 15 s
Calibrated; scales with t
Dark/background rate
0.12 count ch⁻¹ s⁻¹
Conservative effective value
Read noise
0.8 count rms / read
Plausible with EM gain
Excess noise factor
1.0
Photon-counting idealisation
Baseline offset
0 count
Disabled
Shot noise
Poisson(signal + dark)
Active
Digitisation
Round; clip at zero
Active
09
Notes: The exact camera and laser models used experimentally are not yet documented in the repository. These values are defensible for simulation, but hardware datasheets should replace analogues.
[Sources]
- simulator/core/instrument.py
- simulator/core/spectrum.py
- simulator/README.md
- Oxford Instruments iXon Ultra specifications: https://andor.oxinst.com/downloads/uploads/andor-ixon-ultra-emccd-specifications.pdf
- Oxford Instruments EMCCD microscopy specifications (<1 e− with EM gain): https://andor.oxinst.com/products/ixon-emccd-cameras-for-life-science
- HighFinesse/TOPTICA LWA-100k: https://www.toptica.com/fileadmin/Editors_English/06_company/01_news/02_press/2019/20191206_HighFinesse_100k/LWA-100k_EN.pdf

## Slide 10
WHAT “REALISTIC” MEANS HERE
The model is physically plausible; the instrument-specific audit is still required
SUPPORTED
Local frequency axis
Water / PMMA BFS
Peak-count calibration
Exposure scaling
Poisson statistics
PLAUSIBLE
300 K temperature
100 kHz laser linewidth
0.8-count read noise
0.6 μm PSF
Voigt response
NOT YET VALIDATED
Actual camera model
Measured dark/background rate
Measured laser drift
Instrument response split
720 nm wavelength
Required before claiming a real exposure limit
Record hardware specifications and dark frames; measure repeated-reference BFS drift; refit the simulator; then repeat the threshold sweep with replicates.
10
Notes: This is the defensible conclusion: realistic enough to test algorithms and design calibration experiments, not yet sufficient to claim the instrument's true minimum exposure.
[Sources]
- simulator/README.md: calibration tiers and caveats
- Oxford Instruments iXon Ultra specifications: https://andor.oxinst.com/downloads/uploads/andor-ixon-ultra-emccd-specifications.pdf
- HighFinesse/TOPTICA LWA-100k: https://www.toptica.com/fileadmin/Editors_English/06_company/01_news/02_press/2019/20191206_HighFinesse_100k/LWA-100k_EN.pdf